Using a Gas Chromatograph:
Identifying Unknown Compounds

There are many different types of chromatography: paper, thin layer (TLC), liquid, high-pressure liquid (HPLC), and gas (GC). Chromatography is applied in many fields. Biochemists use liquid chromatography to separate proteins; chemists use GC, TLC, and HPLC to identify organic compounds. Forensic scientists and other specialties use gas chromatography for drug tests, toxin screens, and environmental analysis.

All types of chromatography use the same principles that include a stationary phase and a mobile phase. The stationary phase is immobile on the column or the plate and the mobile phase travels from a start point to an end point. Compounds travel from the start to the end at a specific rate depending on their competing affinity for the mobile gas/liquid phase versus the stationary solid phase. Compounds adhere to the stationary phase through dipole interactions, dispersion forces, or ionic interactions.

The Vernier Mini GC uses a metal outer column with a coated stationary phase. A sample, consisting of one or more compounds, is injected into the column and is carried by atmospheric air, which is the mobile phase. Organic compounds flowing out of the chromatography column are indicated as a peak on a chromatograph, as shown in Figure 1. The unique time it takes for a specific compound to exit the column after it is injected is called the retention time. Using a GC, a compound can be identified from a mixture of compounds by its retention time.

Figure 1  Sample gas chromatogram

Several factors can affect a compound’s retention time. More volatile compounds (i.e., compounds with a lower boiling point) will move through the column faster because they are flowing in the mobile phase and not strongly bonded with the stationary phase. The surface functional groups present on the compound are also a factor. For example, alcohols may weakly bond with a polar stationary phase more than esters because alcohols are capable of forming hydrogen bonds. The molecular weight of a compound may also play a role to a slight extent, although it is not a direct relationship between molecular mass and the time it will travel through a GC column.

In this experiment you will explore the process of identifying one or more unknown species using gas chromatography. First, you will practice using a gas chromatograph by testing several known substances. You will then use this information to identify the substances present in an unknown mixture.

OBJECTIVES

In this experiment, you will

- Measure and analyze the retention time of each of six substances as they pass through a Vernier Mini GC.

- Measure and analyze the retention time of an unknown mixture of the substances.

- Identify the substances present in the unknown mixture based on retention times.

MATERIALS

| MacBook | ethanol |
| --- | --- |
| Logger Pro | propyl acetate |
| Vernier Mini GC | butyl acetate |
| 1 (L glass syringe | 2-butanone |
| Kimwipes® | 4-methyl-2-pentanone |
| methanol | unknown mixture |

PRE-LAB EXERCISE

Complete the table below. This information is a common starting point for understanding the behavior of a set of substances, being testing by gas chromatography, which may be found in a mixture. For each substance, identify it by its organic family: alkane, alkene, alcohol, aldehyde, ester, ether, ketone, etc.

| Compound | Boiling temperature ((C) | Molar mass (g/mol) | Bonding functional group |
| --- | --- | --- | --- |
| propyl acetate |  |  |  |
| butyl acetate |  |  |  |
| methanol |  |  |  |
| ethanol |  |  |  |
| 2-butanone |  |  |  |
| 4-methyl-2-pentanone |  |  |  |
| acetone |  |  |  |

PROCEDURE

Part I - Test Known Compounds

1.	Obtain and wear goggles. Gather your materials and equipment and bring to your lab station.  Pour approximately 10mL methanol into a 50mL beaker to keep at your station for cleaning between trials.

2.	Obtain a glass syringe and a set of vials containing the six known substances and one unknown mixture containing three substances to be identified. You will not only test methanol but use it to clean the syringe needle.

3.	Prepare the Vernier Mini GC for data collection.

a.	Turn on and log onto your MacBook computer (username: period(#); password (#) )

b.  Turn on the Mini GC.

c.	Connect the Mini GC to the USB port on your computer or LabQuest.

d.	Open LoggerPro (icon at bottom with graph paper and caliper) and then choose “New” from the File menu

e.	To bring up the Temperature-Pressure profile, tap ►

f.	Set the Temperature-Pressure values to:

| Start temperature | 35(C |
| --- | --- |
| Hold time | 1 min |
| Ramp rate | 10(C/min |
| Final temperature | 65(C |
| Hold time | 6 min |
| Total length | 10.0 min |
| Pressure | 4.0 kPa |

f.	Select Done to initiate the Mini GC warm up. When the Mini GC is ready for injection in Step 7, the message will read, “Inject and select Collect simultaneously”, and the LED will turn to green. Continue with Step 4 during warm up.

4.	Follow the steps below to clean and flush the syringe with methanol. Important: The glass syringe is fragile. Be careful not to bend the needle or bend the plunger. Never pull the plunger back more than 50% of its total volume. Be careful not to bend the plunger as you press it down.

a.	Depress the plunger fully.

b.	Submerge the tip of the syringe needle into the vial of methanol.

c.	Pull back the plunger to fill the barrel about 1/3 full of methanol. Examine the barrel of the syringe and estimate the amount of methanol in the barrel.

d.	Expel the liquid onto a Kimwipe or a paper towel.

e.	Repeat Steps a–d at least two times, until you are comfortable pulling up a liquid into the syringe and measuring the volume in the syringe barrel. Use a Kimwipe® or a paper towel to carefully pat around the tip of the syringe needle.

5.	Follow the process in Step 4 to clean and flush the syringe with propyl acetate.

6.	Collect a volume of propyl acetate for injection.

a.	Submerge the tip of the syringe needle into the vial of propyl acetate one last time.

b.	Draw up 0.10 (L of liquid. It is important that the volume be very close to 0.10 (L and that you inject the same volume for each test.

c.	After collecting your sample, gently wipe the needle, from barrel to tip, with a lint-free tissue such as a Kimwipe.

7.	Prepare for injection and the start of data collection.

a.	The Mini GC should now have reached the correct start temperature and pressure and the LED should be green.

b.	To insert the needle of the syringe into the injection port of the Mini GC, hold the syringe with one hand and steady the needle with your other hand, as shown in Figure 2. Insert the needle into the injection port until the needle stop is fully seated. If the needle sticks, rotate it slightly while inserting. Do not move the plunger yet.

c.	Simultaneously, depress the syringe plunger and select Collect (►) to begin data collection. Pull the needle out of the injection port immediately.

8.	While the data collection proceeds, repeat Step 4 to thoroughly clean the syringe and needle. It may take more than three flushes to feel the syringe plunger move smoothly again, which is your indicator that the syringe and needle are both suitably clean.

9.	The data collection will end after 10 minutes. You may stop the data collection early if you are certain that the entire injected sample has passed through the detector (the line is flat for 2 min after the peak).

10.	Determine the retention time for your chromatogram.

- Choose “Analyze” from the top ribbon of your computer screen, then choose “advanced” then “peak integration” then “signal”.

- Select and integrate the left-most peak. To do this, drag from a little before the peak to a point far enough to the right so that the entire peak is included. Choose Add.

- Record the displayed retention time in your data table.

- Click Cancel to return to the graph.

11.	Label and store your run of data.

- Double click latest, which is the title of your run of data in the table, to the left of the graph.  In the Data Set Options dialog box, type in the name of your compound.  Click OK.  Choose “Store Latest Run” from the experiment menu to store your chromatogram.

12.	Follow the steps below to test the next sample in your set of substances.

a.	Click Collect in Logger Pro, ( ►) to bring up the Temperature-Pressure profile again. This profile will be the same as for your previous run. Select Done to initiate the Mini GC profile.

b.	While the Mini GC adjusts to its Temperature-Pressure profile, repeat Steps 5 and 6 with the next sample.

c.	After the Mini GC is ready, repeat Steps 7–11 to test your new sample.

13.	Repeat Steps 5–11 for the remaining four known substances.

Part II  Test the Unknown Mixture

14.	Clear the graph of data before testing your unknown.

- In Logger Pro, choose Graph Options from the Options menu. Choose Axes Options. Scroll down the list in the y-axis box and remove the check marks in each of the runs for your known compounds. Leave only the Latest box checked. Click done.

15.	Repeat Step 5–11 to test your unknown mixture, but this time inject 0.3 (L and do not store it. When the run is completed, change its name to Unknown.

16.	After you have completed the test of your unknown mixture, save your file for use at a later time. Save the file to the desktop with a name that you will be able to find later.

17.	Turn off the Mini GC. Keep your test results open in Logger Pro or in LabQuest App; you will need to refer to the various chromatograms to answer the Data Analysis questions.

DATA TABLE

| Compound | Retention time (min) |
| --- | --- |
| 1: butyl acetate |  |
| 2: 2-butanone |  |
| 3: 4-methyl-2pentanone |  |
| 4: Acetone |  |
| 5: Propyl acetate |  |
| 6: Ethanol |  |
| 7: Methanol |  |
| 8: Unknown Component 1: Component 2: Component 3: |  |

DATA ANALYSIS

- Discuss the retention times of the six substances with regard to their molecular weights. Describe any pattern that emerges.

- Discuss the retention times of the six substances with regard to their boiling points. Describe any pattern that emerges.

- Identify the substances that are present in your unknown mixture. Support your identification.

# Text boxes

Figure 2